## All
No. | Product Line | Question | Answer
1 | NVR | When the thermal camera is linked to the NVR, after the alarm pops up, the warning sound is delayed or does not sound | After repeated comparison tests, it was found that after the February version and the early March version were upgraded to the 0314 version, the front-end and back-end initialization problems can be solved At present, most cameras‘ problem can be solved, and a small part suspects that the NVR firmware vision is too old and incompatible. The NVR side will offer new firmware in late April. Firstly, upgarde the camera to 0314, then restart& restore factory. if problem not solved, wait for the new version of the baseline.
2 | Camera | The baseline 1217 camera detects the target and the temperature continues to drop. (Baseline thermal camera is used for human body temperature detection) | The problem is known, the version is defective, there is currently no resource correction. There is a temporary solution --- ask zhanhui5, manual tool operation method, change the response rate on the camera side.
3 | Camera | 1217/2637 / 2636B unable to store video Event recording function is not done well. (Schedule--event) | All models and versions of thermal cameras have this problem. The problem is recurring at the headquarters. The headquarters has an improvement plan, but the schedule is ominous -Compatibility issues, currently not supported -Currently only supports all day mode recording to manually select event playback
4 | 4200 | 3.2.0.10 version ivms-4200 export thermal alarm screenshot only thermal channel alarm screen, no optical channel screen | Feedback Zhan Hui, get the latest 4200 version 3.2.0.50 (ftp). The official website is still the old version.
5 | 4200 | The latest version of the baseline 4200 is connected to the thermal camera Event search. When exporting the temperature measurement alarm event pictures, only the thermal channel pictures are available. | Upgraded special temperature measurement version 4200 (supported after V3.2.0.50)
6 | Camera | thermal latest version 5.5.56 V200330 temperature measurement is not accurate | The distance is set to fix, adding blackbody, turning off the face capture. If problem can not be optimized, finally try downgrade.
7 | Handheld | thermal problem: 21B cannot connect to WIFI (both office and mobile phones are not allowed). Enter password-connection failed. There is currently only one device. Restart can not be resolved. | try use 21B as AP mode, let the computer connect.
8 | Camera | The thermal channel of 2TD2636B-15 will become a white screen after working for a period of time, you can see the temperature logo, but you can't see the image, you can only maintain it for a while after restarting and replugging the network cable. | 20% of the bold body is placed in the detection frame, after the customer is removed, everything is normal now.
9 | Handheld | Is only TP21B handheld temperature measurement in human body temperature measurement products supporting Hik-Thermal? | yes
10 | NVR | Does I series NVR and Super Brain need to add the device to the 4200 if they want to receive the human body temperature alarm message? | No, just do it locally. However, it should be noted that now that the human body temperature measurement camera is added to the I series NVR, it may happen that only one channel is added, which can be added on PCweb
11 | HCP | How about compatibility with HikCentral? Can the temperature measurement analysis report be generated? | HCP has just released a customized version for human body temperature measuring, which supports general functions such as preview alarm, but has not yet been released.
12 | other | The usage scene of the device is used in entrance, but we recommend that people enter the scene from the outside, and check it after 5 minutes. In this case, is it impossible to recommend this solution at the entrances of some shopping malls, airports and other scenes? | If the temperature difference between indoor and outdoor is not too large and the impact is small, it can be used. It is recommended to use a solution that guides people to spend more time indoors before entering the temperature measurement environment (preferably 3-5 minutes); If the temperature difference is more obvious and the temperature is relatively low / high (no wind in the room), manual compensation can be considered for the bullet and turret, but there is no relevant setting for handheld.
13 | other | Can hik-thermal app receive temperature alarm? | No, no event center in hik-thermal (in development plan)
14 | Handheld | Can TP21B be added to hik-connect？ | Hikconnect can support live-view after being added, but there is no event alarm till now. (in the development plan)
15 | Handheld | Can TP21B work with black body together? | "Yes, but the function is limited. There are currently two ways to use: ① In Auto mode, that is, after fixing the relative position, place the black body on the screen for real-time calibration (need to draw the rule box), but the logic of the linkage alarm with the 4200 is more confusing when there is a rule box (whether the alarm is only Depends on the highest temperature of the screen, and has nothing to do with the rule frame). This solution is not recommended at this time, and may be optimized in the future. ②Manual mode is equal to do a one-time calibration by using the blackbody. Within thirty minutes after the calibration, the measurement accuracy is increased from the original ± 0.5 ℃ to ± 0.3 ℃. "
16 | Handheld | Can TP31B support sound alarm? | No
17 | Handheld | What are the temperature measurement units of TP21B and TP31B? | Celsius (° C), Fahrenheit (° F), Kelvin (K)
18 | NVR | Are non-I series NVRs compatible with thermal equipment? For storage only. | Yes, you can see the temperature frame information in the storage video screen when you set the stream overlay.
19 | other | Can human body temperature measurement thermal product receive alarm through ONVIF | Yes.
20 | other | Does Milestone support the reception of temperature alarms from body temperature measuring cameras | Yes. The latest plugin is already supported.
21 | Handheld | Can the TP31B be connected externally or projected to the monitor? | No, it does not support access to all functional networks
22 | Handheld | How to display TP21B interface on the monitor? | TP21B is added to the VMS 4200 through WIFI, and the computer video output interface is connected to the monitor
23 | Handheld | Can the TP21B be directly wired to an external monitor? | Cannot
24 | Camera | Which devices support audible and visual alarms? | The device with the model ending in PA.
25 | Camera | Can PA devices that support sound and light alarms choose an alarm language other than English? | Yes，Event >> Basic Event >> Audible Alarm Output >> Alarm Sound Type
26 | Camera | For PA devices that support sound and light alarms, can the sound be adjusted? | Yes，Event >> Basic Event >> Audible Alarm Output >> Sound Volume
27 | Camera | Do PA devices that support sound and light alarms support custom alarm voices? | Not supported temporarily, in development planning
28 | Handheld | Can the handheld be connected to NVR to store video? | Yes, handheld camera access to the wifi of the network segment where nvr is located
29 | Black Body | Do I need an adapter plate to install the blackbody on the tripod? | No, there is a standard tripod screw hole at the bottom
30 | Camera | What is the function of alarm number setting? | Alarm voice broadcast repeat times
31 | 4200 | The temperature measurement Event picture storage and the selected local disk will generate the iVMS-4200alarmPicture folder, the generated file cannot be opened? | This file is a package of alarm pictures and cannot be opened directly. It needs to be exported in 4200
32 | Black Body | How to connect blackbody and camera? | No need to connect, the blackbody can be used after power on, as a standard temperature source.
33 | Camera | How does resolution affect temperature measurement? | No effect, only 384*288 resolution with long focal length equipment
34 | Camera | What are the main differences between 3mm, 6mm, 10mm, 13mm and 15mm temperature measurement products? | Affect the detection distance and support installation height. No difference in temperature measurement performance
35 | Camera | Failed to save the picture / screen recording on the web to the local computer? | Open the browser in administrator mode and try again
36 | NVR | Does the i-series nvr currently support pass-through alarms? | Support transparent transmission, the current alarm and the previous industrial thermal belong to the same temperature alarm category. The human body temperature measurement alarm, including the face snapshot and temperature data, requires the i-series NVR baseline version released at the end of April
37 | Camera | How to solve the problem that the thermal picture of fixed equipment has double image? | Due to the parallax caused by the near visible lens and the thermal lens, you can turn off the two-light fusion in PIP: Image >> Picture in Picture >> Camera 02 >> Picture in Picture Mode
38 | Handheld | Do I need an adapter plate to mount the handheld device on a tripod? | No, there is a standard tripod screw hole at the bottom
39 | Handheld | Can the TP21B be unplugged the battery and connected to the charging cable? | Yes, but not recommended. Bad operation will cause frequent power outages
40 | Camera | The screen is full screen or partly red, how to close | Image >> Display Settings >> Camera 02 >> Image Enhancement >> Palettes, turn the above enable off
41 | Camera | How to change the pseudo color mode of fixed equipment? | Image >> Display Settings >> Camera 02 >> Image Enhancement >> Palettes
42 | Camera | How to change the temperature font size in the screen of fixed device? | Image >> VCA Rule Display >> Camera 02 >> Font Size, the font color of different temperature ranges can also be adjusted here
43 | Handheld | What does Support audio intercom in TP21B spec mean? | Supports two-way intercom to connect to the event center (including Hik Thermal)
44 | Camera | Does the device need to warm up after plugging and unplugging the power supply for a short period of time after upgrading, restarting, and resetting? | It is recommended to wait for 10 minutes before measuring the temperature if the device only unplugged for a very short period.
45 | Camera | The temperature of the equipment changes slightly after half an hour | After the warm-up time reaches 1 hour, see if it is stable. The warm-up time is related to the actual environment. Generally, most environments are stable enough after 1 hour (based on the Start Time on SADP)
46 | Black Body | After turning on blackbody compensation, the measured temperature of the human body has risen significantly? | If the heat engine is guaranteed to be completed, it may be that the black body temperature control element is affected during transportation, resulting in a true temperature change (the accuracy will not change). The VCA resource mode of the thermal device can be switched to temperature measurement, and the true temperature of the black body in the field of view of the camera can be measured. If the true temperature of the blackbody is 39.3 ° C, you can switch back to the human body temperature measurement mode and fill in the blackbody temperature at the blackbody compensation position as 39.3 ° C.
47 | Black Body | Can the temperature be measured outdoors after using a blackbody? | No, the outdoor environment affects the temperature of the measured skin surface. It has nothing to do with the black body
48 | DeepinMind | There is no temperature frame and temperature value information on the face image captured by DeepinMind | Face Capture >> Overlay & Capture>>enable Display Target Info. on Alarm Picture
49 | DeepinMind | The same person of the DeepinMind conducted multiple face captures and temperature measurements | Face Capture >> Advanced Configuration See if Capture Times is Limited as 1.
50 | Camera | Only the single thermal camera is connected to the computer. Does the current version support over-temperature event storage? The high temperature alarm recording is detected and stored in the 4200 storage server or SD card | Not supported, currently only supports All day recording, manually click to view and playback according to the alarm event.
51 | 4200 | 4200 can only store thermal channel pictures | use 4200 V3.2.0.102
52 | Camera | PA device cannot flash alarm. | Confirm that the temperature measurement configuration linkage and Event >> Basic Event >> Flashing Alarm Light Output are all configured correctly, the function is turned off by default
53 | Camera | Is there a difference between the maximum number of faces detected by the bullet and turret? | No, both support 30 faces.
54 | Camera | Is the firmware of the bullet and the turret the same | Yes
55 | Camera | How to turn off red/green grid in the live view. | Event>>Basic Event>>Motion Detection, colose both channels' relevant function.
56 | Camera | How to reverse the picture when the device is flipped upside down? Will it affect face detection? | Image >> Display Settings >> Camera 0x >> Video Adjustment >>Mirror Up/Down; No effect
57 | Camera | Why can the resolution of the thermal channel in Video be adjusted to 1280 * 720 and others? | This just stretches the pixels of the thermal channel. The resolution of the image generated by the thermal camera does not change.The OSD and regular font size are also stretched.
58 | Camera | The recognition frame does not appear on the screen, but the face detection range frame still display. | Check the interpupillary filter setting, whether the face is filtered out.
59 | Camera | There are no rules, face recognition frame, temperature information in the picture | Verify that Display VCA Info By is selected as Video. Check whether Local has all relevant rules enabled.
60 | Camera | Whether the device can be downgraded | The major version cannot be downgraded, such as 5.5.26_200330 cannot be downgraded to 5.5.24_200220, but it can be downgraded to 5.5.26_200317
61 | 4200 | The 4200 body temperature measurement version has no relevant interface | When installing the software, choose the plug-in AI DASHBOARD to install together
62 | 4200 | Human body temperature version 4200 AI DASHBOARD face recognition speed is very slow | By default, the face capture and temperature information will be uploaded when the face leaves the detection range
63 | Camera | What's new in 5.5.26_200330 compared to 5.5.26_200317 | Add adaptive distance mode; PA device adds multilingual alarm voice; default forehead temperature measurement; other specific can see the Read me document in the corresponding path in FTP
64 | Camera | Version 5.5.26_200330, how to solve the abnormal temperature measurement on the side face | Generally caused by wrong judgment of the distance when facing the side, if it is not guaranteed that the front of the face is always facing the camera, it is recommended to turn on the fixed distance mode
65 | NVR | Add the thermal camera directly to the i series NVR at the NVR side, only one channel is added successfully | Log in to the IP of the i-series NVR from the computer and add the device from the WEB page to solve this issue.
66 | Camera | The bullet can be directly installed on the tripod with standard screw holes. Does it rely on the adapter plate? | No need the adapter plate, after the screw on the top of the bullet is unscrewed, the protective cover can be directly flipped after removing the cover, and then the upper / lower mirror can be used in the configuration. But this not looks very nice, and the angle cannot be adjusted. It can only guarantee the level of sight and the same height as the top of the tripod. This way is not recommended.
67 | other | Why the device cannot be added to Hik-Connect | Network >> Advanced Settings >> Platform Access>>configure Hik-connect
68 | other | Why the human body temperature measuring device added to Hik-Connect cannot receive the alarm | It is recommended to add it by scanning the QR code or filling in the serial number; adding directly via IP will not receive the alarm
69 | Camera | What are the Camera's Audio in and Audio out interfaces for? | Audio Out is a 3.5 mm headphone jack. For P devices, it can only be used as the earphone end in two-way intercom. For the PA device, it can not only be the earphone end of the two-way intercom, but also can transmit the original alarm voice of the PA device.
70 | Camera | How to use ALARM OUT interface? | Alarm Out is equivalent to a I/O signal. After connecting an external active alarm device (such as a buzzer), the switch closes after the device alarms, and the external alarm will alarm. This applies regardless of whether the device is P / PA. Note that Alarm Out cannot transmit audio signals.
71 | Camera | Can the ALARM OUT interface transmit audio signals? | cannot
72 | Handheld | What does Offline Picture and Thermal Graphic mean in the camera setting of handheld products? | 1. Thermal Graphic is to save a snapshot; 2. Offline Picture can support secondary analysis, such as redrawing the area for temperature measurement, etc.
73 | Camera | Does thermal camera voice broadcast support Myanmar language? | not support. The latest version will be released at the end of April in the future to support custom voice entry settings
74 | Camera | Can the temperature highlight function and font size of the thermal camera be changed? | Yes. Image-VCA Rule Display.
75 | DeepinMind | How to export the report in DeepinMind | 1. Only the information of Check in and Attandacne in DeepinMind can be exported in Excel format; 2. Other alarm information or personnel search does not support Excel format export, only photos and corresponding code information can be exported, but the exported photos contain temperature information;
76 | Camera | When does the Upload Captured face Image function of the thermal camera need to be opened? Is it used to automatically upload captured pictures? | 1. The upload to the alarm center is enabled by default. If Upload Captured face Image is not enabled, the platform or 4200 can only receive alarm information, and cannot receive photos of face capture pictures; 2. Therefore, when the thermal camera works alone, that is, when there is no platform, in order to reduce the pressure on the bandwidth, it is recommended not to enable Upload Captured Face Image; if there are terminals or platforms such as 4200 or DeepinMind or HCP, you must check Upload Captured Face Image
77 | Camera | Camera upgrade failed | 1. Clear the browser cache; 2. Download the upgrade file again; 3. Restore camera after re-upgrade;
78 | Camera | If there is a temperature alarm on the screen, if the person with high temperature does not leave the screen, there will be no alarm | 1. The camera uses quick shot mode, and there is no situation that people do not leave the screen and do not take pictures; 2. It is recommended to check the parameter configuration in quick shot to restore the threshold to the default 0. limite is also restored to the default 1
79 | Camera | Is the SDK of the body temperature measuring camera the same as the SDK of the ordinary camera? | The human body temperature measurement camera can also use the SDK package on the official website, but the handheld device does not support
80 | Camera | The handheld product is connected to HIK-thermal, which can record video but cannot capture pictures | At present, HIK-Thermal does have this defect, and the R & D is currently being resolved. The specific time point has not been confirmed, and new progress will be notified in time.
81 | Handheld | Hand-held TP21B, enter the WIFI password, after entering the password, click OK and nothing happens | 1. It is an occasional problem, and there is very little feedback at present, and it is difficult to confirm the reason for R & D; 2. It is recommended to initialize the device and then reconnect; 3. In addition, when the password is entered, choose Close, not Enter. Therefore, consider the password error caused by entering the character "Enter" in the password;
82 | Handheld | 1. Does the handheld support three-party batteries; 2. One charger can charge several batteries at the same time; 3. Can the device be used normally when recharging; | 1. The handheld product supports three parties' batteries; 2. One charger can charge 2 batteries at the same time; 3. The device can be used normally during the charging process;
83 | Handheld | Can the handheld product TP21B be connected to HCP? | After testing, I found: 1. HCP can access TP21B high-end handheld devices; 2. It can be previewed normally, and has been dual-channel with ordinary thermal camera; 3. The recording plan cannot be configured, but you can manually capture pictures and record videos on the Liveview interface; 4. Temperature and alarm can be configured, and the alarm information contains temperature information;
84 | Handheld | Can handheld devices connect to NVR? | 1. The handheld device supports access to Class I NVR; 2. However, it does not support the configuration of recording plans including various events such as event recording, scheduled recording, and 24-hour recording; 3. Only manual picture capture and recording are supported on the live-view interface;
85 | Handheld | Handheld device WIFI hotspot can access several mobile phones | In theory, up to 5 devices can be connected, but it is recommended not to exceed 3 devices.
86 | Handheld | Handheld device connected to Hik-Connect cannot receive high temperature alarm information | 1. The handheld device currently supports access to HIK-Connect; 2. However, Ezviz has not yet completed the high-temperature alarm function and is currently advancing. 3. If you are eager to use this function, you can Baidu [temperature measurement wizard], this app is developed by Ezviz;
87 | other | 1. What do these icons mean? 2. Why the temperature threshold is cleared at the click and the camera keeps alarming | 1. Respectively measure the center temperature, the highest temperature point, the lowest temperature point, and the frame temperature; 2. This is a bug of the current product. The temperature threshold will be cleared when clicked, and the camera will always alarm; 3. Currently solving this defect;
88 | Handheld | Where is the Device Password of the handheld device set? I forgot my password | 1. The device password of the handheld device is set when the device is activated for the first time; 2. If you forget the password, you can only reset the device and set it after re-activation;
89 | Camera | thermal camera connected to 4200, but failed to export photos | Customer's individual computer problem, the problem is solved after replacing the computer;
90 | Camera | Human body temperature 36.4 ~ 36.5 when black body exists; When there is no black body, the temperature is 37.2 ~ 37.5; | 1. The accuracy of the blackbody is affected during transportation; 2. Change the camera to temperature measurement mode and test the blackbody temperature after half an hour of warm-up; 3. Fill in the actual temperature of the black body after measurement into the black body temperature value at the WEB end;
91 | Camera | Human body temperature is too high | 1. After careful inquiry, it is found that the temperature of other people is normal, and the temperature of only 2 people is high; 2. This is an individual difference. Not everyone ’s skin emissivity is 0.98. Individual differences are inevitable; at present this problem cannot be solved
92 | HCP | HCP + 2TD1217B, when there is no DeepinMind, will there be real-time temperature records here? | 1. The Body Temperature module is intended for temporary scenarios. In this scenario, there is no DeepinMind to participate in; 2. This interface will display the real-time temperature; 3. However, this interface is not a Live View interface, it will only display a picture of the person who triggered the threshold; 4. There are two thresholds on the HCP. One threshold is used to filter out extreme cases of abnormally low body temperature (such as fake faces), and the other threshold is used to filter out abnormal body temperatures (such as fever);
93 | HCP | After setting the temperature on the HCP temperature measurement module Body Temperature, only the people with abnormal high temperature will be displayed, and the people with normal body temperature will no longer be displayed. | 1. There are two thresholds on the HCP. One threshold (Thermometry Threshold) is used to filter out extreme cases of abnormally low body temperature (such as fake faces), and the other threshold (Alarm Threshold) is used to filter out abnormalities (such as fever); 2. Thermometry Threhold will be synchronized to the alarm value and pre-alarm value of the device after the setting is completed, and the values of these two parameters remain the same; 3. Set the Thermometry threshold to about 35, which is used to choose less extreme low-temperature conditions; set the Alarm Threshold to the high-temperature alarm threshold, which will display the normal body temperature and the body temperature of people with abnormal temperature;
94 | Black Body | 1. An individual error occurs during measurement. The temperature measurement of other people is normal, but the temperature measurement of individual people is abnormal; 2. This situation is weakened when there is a black body, and it is more obvious when there is no black body; | 1. We measure the temperature of a person's forehead and the infrared radiation emitted by the person's forehead; 2. Therefore, not everyone's forehead temperature and body temperature are the same; at the same time, not everyone's emissivity is 0.98; 3. The black body can reduce the error, so when the black body exists, it can reduce the impact of this individual difference; 4. If such individual differences often occur, and customers are more concerned about this problem, it is recommended to deploy a black body in the system to reduce the occurrence of this problem
95 | DeepinMind | After the upgrade of DeepinMind, there is no captured face picture | Upload captured face Image is not enabled on the camera side, and the fault is resolved after the function is enabled;
96 | Camera | After the Alarm out is connected, after the relevant linkage is configured on the temperature measurement interface, the external alarm does not sound, and neither 1A1B nor 2A2B can alarm | Alarm Schedule in Basic Event is not configured, the alarm is normal after configuration
97 | other | Hikconnect and human body temperature measurement equipment alarm logic | Currently: the same channel, only one event will be uploaded every minute
98 | Camera | Any change for new firmware?(200415) | There is no difference on the user interface of new version (compared to version 200330). This version have a better performance when the device is used in environment with relatively high temperature (like 30℃ in a tent). Also some bugs have been fixed.
99 | Camera | The highest temperature is at the junction of the face and background, resulting in frequent high temperature false alarms | It can be solved after upgrading to the version after 200330
100 | Camera | Now the description of the small language interface still retains English description | The document group did not keep up with the progress of the interface update, and the next major version will be updated to solve the problem
101 | Camera | Whether the visible light channel of the human body temperature measuring camera supports active IR | Supported, but not recommended. Use as much as possible in an environment with sufficient light source
102 | Camera | Hikvision logo still displayed after neutral firmware upgrade | Clear the browser cache and restart to solve
103 | Camera | Human temperature measurement firmware upgrade failed | Confirm that the major version of the upgrade package is after the current version. Open the browser in administrator mode and try to upgrade again
104 | DeepinMind | After upgrading the firmware of 77 DeepinMind, neither arming schdule nor audio warning can be enabled | Device upgrade failed, clear browser cookies, re-upgrade, restore device after upgrade
105 | DeepinMind | There are audio warnings in No mask linkage and abnormal linkage in DeepinMind. What is the difference between this and voice linkgae? | If you want to use the audio warning in no mask linkage and abnormal linkage, you must first enable the relevant options in voice linkage (Mask, Temp Normal, Temp Abnormal)
106 | Camera | The camera alarm output function is unsuccessful, but the alarm output can be performed manually | The sechdule of the alarm output is not configured on the webpage, it is necessary to configure the sechdule
107 | Camera | The camera has a capture, the face is matched, but there is no other structured information of the face | 1. To obtain other structured information, VCA resource needs to select multi-type detection; 2. Face capture is only a face capture, face comparison, but no other structured information;
108 | Camera | The camera still shows high temperature after warming up | Upgrade the firmware version to 0415, V.0415 optimizes the effect of background temperature;
109 | Camera | The OSD time of the thermal channel and the visible light channel are different | This problem does exist at present, R & D is already known, and it will be fixed in the future
110 | Camera | In scenes where no one is present, the camera still tests the maximum ambient temperature in the screen | The camera does not have Enable face detection; therefore the camera detects the highest temperature point in the environment;
111 | Handheld | The picture of thermal channels are all orange | 1. It is necessary to cancel Above in Image Enhancement, that is, to cancel the function of high temperature highlighting; 2. The initial suspicion is that the ambient temperature in Indonesia is relatively high, and the ambient temperature exceeds the threshold temperature, resulting in the high temperature of the entire picture highlighting;
112 | NVR | 1. The latest firmware version 0414 of the I series NVR will be released soon; 2. Human temperature display can be displayed in this version; 3. Sound alarm of temperature status; 4. Search and playback of temperature measurement results; 5. Temperature information transparent transmission alarm; 6. The event recording is still unsuccessful. Even if the event recording is selected, the event is still being recorded, but the NVR actually still records 24 hours; | R & D has been informed and they will improve afterwards
113 | Camera | Abnormal temperature measurement after the camera is taken from the outdoor to the indoor | 1. The camera body is made of iron material, which absorbs heat quickly. The high outdoor temperature leads to the high temperature of the body, which affects the detection of the external temperature of the internal chip of the camera; 2. Take it from outdoor to indoor, after cooling for a period of time, use it after reheating the machine;
114 | Camera | For the camera infrared channel, the human body has ghost images | The infrared channel has turned on dual-light fusion, and needs to be turned off
115 | Handheld | What is the capacity of the SD card that comes with the handheld device when it is shipped? | Economical 8G, high-end 16G
116 | Handheld | 1. In order to save storage space, does the handheld support capture screenshots? 2. Does SD card support storage coverage? 3. Can I use USB to export videos and screenshots? 4. Why is it written in the brochure that can be applied to the transformer substation scene? | 1. Those who do not support event capture; 2. The SD card does not support storage coverage; 3. Support USB export video and snapshot; 4. The customer downloads the TP31 brochure, TP31 and TP31B are different, TP31B measures the temperature of the human body, TP31 is not;
117 | Handheld | TP31 handheld device startup screen is pink | The camera movement is damaged and needs to be replaced free of charge
118 | 4200 | 4200 AI dashboard can't find the place to export data | The 4200 AI dashboard interface does not support exporting data, but you can search by event and export in Event;
119 | 4200 | What are the meanings of P1 and P2 on the TP21B interface | TP21B supports setting the maximum temperature point (Max.), The minimum temperature point (Min.) And the central temperature point (Cen.), And also supports user-defined point temperature measurement, P1 means the first custom point (Position) 1) And up to 10 custom temperature measurement points can be set;
120 | DeepinMind | DeepinMind audible alarm failed | 1. DeepinMind supports audible alarms, but DeepinMind does not have a built-in speaker and requires an external speaker; 2. The camera also supports voice alarm, but only -P's camera comes with a built-in speaker, so no external speaker is required;
121 | Camera | The camera failed to connect to the security gate through the RS485 line | 1. The thermal imaging camera for the security gate is a customized version, and the firmware version of the thermal imaging camera for human body temperature measurement on the FTP is not universal; 2. Reuse the original firmware package to upgrade the camera to a customized version of the security package;
122 | HCP | HCP temperature measurement module Body Temperature, there is no way to link the camera | 1. The HCP temperature measurement module (Body Temperature) is for the temporary deployment of the scene, that is, the thermal imaging camera stand alone; 2. This module does not support indirect association via DeepinMind, and can only be directly connected to HCP;
123 | Camera | For human body temperature cameras, the number of streams for the visible light channel and the thermal channel | Both visible light and thermal imaging channels support main and sub-stream
124 | Camera | How to turn off the sound of the camera itself | Directly turn off the camera's sound alarm linkage or adjust the volume
125 | DeepinMind | The DeepinMind connected to the ordinary camera of the temperature measurement scheme can display the face comparison on the UI interface. Is it detected whether it is with a mask? | Can
126 | Handheld | If the handheld temperature measurement tp21b is connected to nvr, if only the stream is connected to the NVR, is there temperature information superimposed on this stream? Can I see temperature information on the NVR? | Yes, TP21B itself has temperature information superimposed on the video stream.
127 | 4200 | Can the 4200 temperature measurement version offer a Minority language voice package? | Multilingual packages are expected to be given in 4.24
128 | Camera | Does the altitude have an effect on the human body temperature measurement | Barometric pressure has no effect on temperature measurement, but the ultraviolet light is strong in high altitude areas, and it needs to be used in an environment without sunlight exposure / reflection
129 | 4200 | Can the 4200's AI dashboard display face comparison and temperature measurement at the same time? | can
130 | NVR | Whether I series NVR alarm sound can be customized | The baseline does not support modification, it can only trigger the buzzer to sound
131 | NVR | When will the new NVR baseline firmware for body temperature measurement cameras be released | The beta version is already on ftp at / 01 All Materials of Temperature Screening Solution / 02 NVR / 05 Firmware / I Series NVR / The official baseline version will be released before mid-May
132 | Camera | Can the sound and light of PA be turned on and off separately | Yes, configuration in linkage
133 | Camera | camera linkage email only sends a picture of the thermal channel | Solve by upgrading to the latest version
134 | Handheld | Does TP21B high-end handheld support Onvif and ISAPI protocols? | Does not support Onvif, only supports ISAPI and SDK
135 | Handheld | Does the TP21B tripod connected to the standard screw interface need an adapter plate? | No need
136 | NVR | Can the web interface of the I series NVR directly configure the camera's body temperature measurement configuration | Add via IP: not supported; via POE interface of NVR: can be configured with virtual host "
137 | 4200 | What is the latest version of 4200 in FTP? | 3.2.50.2, the latest customized version released for Haikang's full line of human body temperature measurement products
138 | DeepinMind | Is the front-end or back-end do the screenshot in the DeepinMind solution? | The working logic is as follows: the camera visible light channel captures the face + superimposed temperature information-the snapshot is given to the DeepinMind, the DeepinMind performs temperature classification, mask detection and face comparison
139 | NVR | Does the latest firmware of the I series NVR support event filtering based on the measured temperature | can support
140 | Handheld | Is the TP21B handheld add to 4200 by 8000 port? | The default is port 8000
141 | Handheld | What is the capacity of the expansion card supported by the handheld device? What are compatible brands? | Maximum support 128G; minimum configuration: class 4, read / write rate 5M / S; brand: Hikvision, Toshiba, SanDisk / Kingston are preferred, other brands do not guarantee compatibility.
142 | Camera | Does the camera support analog output? | Currently only 2637 models support this function, the local output is directly output by the movement
143 | Handheld | Will the temperature information be automatically displayed after the TP21B is connected to the NVR? | Can
144 | Camera | What is the role of the camera's 485 interface | The terminal (such as security gate) can receive the temperature information transmitted from the 485 port through the modbus protocol

## Camera
No. | Product Line | Question | Answer
1 | Camera | The baseline 1217 camera detects the target and the temperature continues to drop. (Baseline thermal camera is used for human body temperature detection) | The problem is known, the version is defective, there is currently no resource correction. There is a temporary solution --- ask zhanhui5, manual tool operation method, change the response rate on the camera side.
2 | Camera | 1217/2637 / 2636B unable to store video Event recording function is not done well. (Schedule--event) | All models and versions of thermal imaging cameras have this problem. The problem is recurring at the headquarters. The headquarters has an improvement plan, but the schedule is ominous -Compatibility issues, currently not supported -Currently only supports all day mode recording to manually select event playback
3 | Camera | Thermal imaging latest version 5.5.56 V200330 temperature measurement is not accurate | The distance is set to fix, adding blackbody, turning off the face capture can not be optimized, and finally downgrade.
4 | Camera | The thermal imaging channel of 2TD2636B-15 will become a white screen after working for a period of time, you can see the temperature logo, but you can't see the image, you can only maintain it for a while after restarting and replugging the network cable. | 20% of the bold body is placed in the detection frame, after the customer is removed, everything is normal now.
5 | Camera | Which devices support audible and visual alarms? | The device with the model ending in PA.
6 | Camera | Can PA devices that support sound and light alarms choose an alarm language other than English? | Yes，Event >> Basic Event >> Audible Alarm Output >> Alarm Sound Type
7 | Camera | For PA devices that support sound and light alarms, can the sound be adjusted? | Yes，Event >> Basic Event >> Audible Alarm Output >> Sound Volume
8 | Camera | Do PA devices that support sound and light alarms support custom alarm voices? | Not supported temporarily, in development planning
9 | Camera | Do I need an adapter plate to install the blackbody on the tripod? | No, there is a standard tripod screw hole at the bottom
10 | Camera | What is the function of alarm number setting? | Alarm voice broadcast repeat times
11 | Camera | How to connect blackbody and camera? | No need to connect, the blackbody can be used after power on, as a standard temperature source.
12 | Camera | How does resolution affect temperature measurement? | No effect, only 384*288 resolution with long focal length equipment
13 | Camera | What are the main differences between 3mm, 6mm, 10mm, 13mm and 15mm temperature measurement products? | Affect the detection distance and support installation height. No difference in temperature measurement performance
14 | Camera | Failed to save the picture / screen recording on the web to the local computer? | Open the browser in administrator mode and try again
15 | Camera | How to solve the problem that the thermal imaging picture of fixed equipment has double image? | Due to the parallax caused by the near visible lens and the thermal imaging lens, you can turn off the two-light fusion in PIP: Image >> Picture in Picture >> Camera 02 >> Picture in Picture Mode
16 | Camera | The screen is full screen or partly red, how to close | Image >> Display Settings >> Camera 02 >> Image Enhancement >> Palettes, turn the above enable off
17 | Camera | How to change the pseudo color mode of fixed equipment? | Image >> Display Settings >> Camera 02 >> Image Enhancement >> Palettes
18 | Camera | How to change the temperature font size in the screen of fixed device? | Image >> VCA Rule Display >> Camera 02 >> Font Size, the font color of different temperature ranges can also be adjusted here
19 | Camera | Does the device need to warm up after plugging and unplugging the power supply for a short period of time after upgrading, restarting, and resetting? | It is recommended to wait for 10 minutes before measuring the temperature.
20 | Camera | The temperature of the equipment changes slightly after half an hour | After the warm-up time reaches 1 hour, see if it is stable. The warm-up time is related to the actual environment. Generally, most environments are stable enough after 1 hour (based on the Start Time on SADP)
21 | Camera | After turning on blackbody compensation, the measured temperature of the human body has risen significantly? | If the heat engine is guaranteed to be completed, it may be that the black body temperature control element is affected during transportation, resulting in a true temperature change (the accuracy will not change). The VCA resource mode of the thermal imaging device can be switched to temperature measurement, and the true temperature of the black body in the field of view of the camera can be measured. If the true temperature of the blackbody is 39.3 ° C, you can switch back to the human body temperature measurement mode and fill in the blackbody temperature at the blackbody compensation position as 39.3 ° C.
22 | Camera | Can the temperature be measured outdoors after using a blackbody? | No, the outdoor environment affects the temperature of the measured skin surface, and has nothing to do with the black body calibration.
23 | Camera | Only the single thermal camera is connected to the computer. Does the current version support over-temperature event storage? The high temperature alarm recording is detected and stored in the 4200 storage server or SD card | Not supported, currently only supports All day recording, manually click to view and playback according to the alarm event.
24 | Camera | PA device cannot flash alarm. | Confirm that the temperature measurement configuration linkage and Event >> Basic Event >> Flashing Alarm Light Output are all configured correctly, the function is turned off by default
25 | Camera | Is there a difference between the maximum number of faces detected by the bullet and turret? | No, both support 30 faces.
26 | Camera | Is the firmware of the bullet and the turret the same | Yes
27 | Camera | How to turn off red/green grid in the live view. | Event>>Basic Event>>Motion Detection, colose both channels' relevant function.
28 | Camera | How to reverse the picture when the device is flipped upside down? Will it affect face detection? | Image >> Display Settings >> Camera 02 >> Video Adjustment >>Mirror Up/Down; No effect
29 | Camera | Why can the resolution of the thermal imaging channel in Video be adjusted to 1280 * 720 and others? | This just stretches the pixels of the thermal imaging channel. The resolution of the image generated by the thermal camera does not change.The OSD and regular font size are also stretched.
30 | Camera | The recognition frame does not appear on the screen, but the face detection range frame still display. | Check the interpupillary filter setting, whether the face is filtered out.
31 | Camera | There are no rules, face recognition frame, temperature information in the picture | Verify that Display VCA Info By is selected as Video. Check whether Local has all relevant rules enabled.
32 | Camera | Whether the device can be downgraded | The major version cannot be downgraded, such as 5.5.26_200330 cannot be downgraded to 5.5.24_200220, but it can be downgraded to 5.5.26_200317
33 | Camera | What's new in 5.5.26_200330 compared to 5.5.26_200317 | Add adaptive distance temperature measurement mode; PA device adds multilingual alarm voice; default forehead temperature measurement; other specific can see the Read me document in the corresponding path in FTP
34 | Camera | Version 5.5.26_200330, how to solve the abnormal temperature measurement on the side face | Generally caused by wrong judgment of the distance when facing the side, if it is not guaranteed that the front of the face is always facing the camera, it is recommended to turn on the fixed distance mode
35 | Camera | The bullet can be directly installed on the tripod with standard screw holes, don't you rely on the adapter plate? | Yes, after the screw on the top of the bullet is unscrewed, the protective cover can be directly flipped after removing the cover, and then the upper / lower mirror can be used in the configuration. But this not looks very nice, and the angle cannot be adjusted. It can only guarantee the level of sight and the same height as the top of the tripod. This way is not recommended.
36 | Camera | What are the Camera's Audio in and Audio out interfaces for? | Audio Out is a 3.5 mm headphone jack. For P devices, it can only be used as the earphone end in two-way intercom. For the PA device, it can not only be the earphone end of the two-way intercom, but also can transmit the original alarm voice of the PA device.
37 | Camera | How to use ALARM OUT interface? | Alarm Out is equivalent to a I/O signal. After connecting an external active alarm (such as a buzzer), the switch closes after the device alarms, and the external alarm will alarm. This applies regardless of whether the device is P / PA. Note that Alarm Out cannot transmit audio signals.
38 | Camera | Can the ALARM OUT interface transmit audio signals? | cannot
39 | Camera | Does thermal camera voice broadcast support Myanmar language? | not support. The latest version will be released at the end of April in the future to support custom voice entry settings
40 | Camera | Can the temperature highlight function and font size of the thermal camera be changed? | Yes. Image-VCA Rule Display.
41 | Camera | When does the Upload Captured face Image function of the thermal camera need to be opened? Is it used to automatically upload captured pictures? | 1. The upload to the alarm center is enabled by default. If Upload Captured face Image is not enabled, the platform or 4200 can only receive alarm information, and cannot receive photos of face capture pictures; 2. Therefore, when the thermal camera works alone, that is, when there is no platform, in order to reduce the pressure on the bandwidth, it is recommended not to enable Upload Captured Face Image; if there are terminals or platforms such as 4200 or DeepinMind or HCP, you must check Upload Captured Face Image
42 | Camera | Camera upgrade failed | 1. Clear the browser cache; 2. Download the upgrade file again; 3. Restore camera after re-upgrade;
43 | Camera | If there is a temperature alarm on the screen, if the person with high temperature does not leave the screen, there will be no alarm | 1. The camera uses quick shot mode, and there is no situation that people do not leave the screen and do not take pictures; 2. It is recommended to check the parameter configuration in quick shot to restore the threshold to the default 0. limite is also restored to the default 1
44 | Camera | Is the SDK of the body temperature measuring camera the same as the SDK of the ordinary camera? | The human body temperature measurement camera can also use the SDK package on the official website, but the handheld device does not support
45 | Camera | The handheld product is connected to HIK-thermal, which can record video but cannot capture pictures | At present, HIK-Thermal does have this defect, and the R & D is currently being resolved. The specific time point has not been confirmed, and new progress will be notified in time.
46 | Camera | thermal camera connected to 4200, but failed to export photos | Customer's individual computer problem, the problem is solved after replacing the computer;
47 | Camera | Human body temperature 36.4 ~ 36.5 when black body exists; When there is no black body, the temperature is 37.2 ~ 37.5; | 1. The accuracy of the blackbody is affected during transportation; 2. Change the camera to temperature measurement mode and test the blackbody temperature after half an hour of warm-up; 3. Fill in the actual temperature of the black body after measurement into the black body temperature value at the WEB end;
48 | Camera | Human body temperature is too high | 1. After careful inquiry, it is found that the temperature of other people is normal, and the temperature of only 2 people is high; 2. This is an individual difference. Not everyone ’s skin emissivity is 0.98. Individual differences are inevitable; at present this problem cannot be solved
49 | Camera | After the Alarm out is connected, after the relevant linkage is configured on the temperature measurement interface, the external alarm does not sound, and neither 1A1B nor 2A2B can alarm | Alarm Schedule in Basic Event is not configured, the alarm is normal after configuration
50 | Camera | Any change for new firmware?(200415) | There is no difference on the user interface of new version (compared to version 200330). This version have a better performance when the device is used in environment with relatively high temperature (like 30℃ in a tent). Also some bugs have been fixed.
51 | Camera | The highest temperature is at the junction of the face and background, resulting in frequent high temperature false alarms | It can be solved after upgrading to the version after 200330
52 | Camera | Now the description of the small language interface still retains English description | The document group did not keep up with the progress of the interface update, and the next major version will be updated to solve the problem
53 | Camera | Whether the visible light channel of the human body temperature measuring camera supports active IR | Supported, but not recommended. Use as much as possible in an environment with sufficient light source
54 | Camera | Hikvision logo still displayed after neutral firmware upgrade | Clear the browser cache and restart to solve
55 | Camera | Human temperature measurement firmware upgrade failed | Confirm that the major version of the upgrade package is after the current version. Open the browser in administrator mode and try to upgrade again
56 | Camera | The camera alarm output function is unsuccessful, but the alarm output can be performed manually | The sechdule of the alarm output is not configured on the webpage, it is necessary to configure the sechdule
57 | Camera | The camera has a capture, the face is matched, but there is no other structured information of the face | 1. To obtain other structured information, VCA resource needs to select multi-type detection; 2. Face capture is only a face capture, face comparison, but no other structured information;
58 | Camera | The camera still shows high temperature after warming up | Upgrade the firmware version to 0415, V.0415 optimizes the effect of background temperature;
59 | Camera | The OSD time of the thermal channel and the visible light channel are different | This problem does exist at present, R & D is already known, and it will be fixed in the future
60 | Camera | In scenes where no one is present, the camera still tests the maximum ambient temperature in the screen | The camera does not have Enable face detection; therefore the camera detects the highest temperature point in the environment;
61 | Camera | Abnormal temperature measurement after the camera is taken from the outdoor to the indoor | 1. The camera body is made of iron material, which absorbs heat quickly. The high outdoor temperature leads to the high temperature of the body, which affects the detection of the external temperature of the internal chip of the camera; 2. Take it from outdoor to indoor, after cooling for a period of time, use it after reheating the machine;
62 | Camera | For the camera infrared channel, the human body has ghost images | The infrared channel has turned on dual-light fusion, and needs to be turned off
63 | Camera | The camera failed to connect to the security gate through the RS485 line | 1. The thermal imaging camera for the security gate is a customized version, and the firmware version of the thermal imaging camera for human body temperature measurement on the FTP is not universal; 2. Reuse the original firmware package to upgrade the camera to a customized version of the security package;
64 | Camera | For human body temperature cameras, the number of streams for the visible light channel and the thermal channel | Both visible light and thermal imaging channels support main and sub-stream
65 | Camera | How to turn off the sound of the camera itself | Directly turn off the camera's sound alarm linkage or adjust the volume
66 | Camera | Does the altitude have an effect on the human body temperature measurement | Barometric pressure has no effect on temperature measurement, but the ultraviolet light is strong in high altitude areas, and it needs to be used in an environment without sunlight exposure / reflection
67 | Camera | Can the sound and light of PA be turned on and off separately | Yes, configuration in linkage
68 | Camera | camera linkage email only sends a picture of the thermal channel | Solve by upgrading to the latest version
69 | Camera | Does the camera support analog output? | Currently only 2637 models support this function, the local output is directly output by the movement
70 | Camera | What is the role of the camera's 485 interface | The terminal (such as security gate) can receive the temperature information transmitted from the 485 port through the modbus protocol

## Black Body
No. | Product Line | Question | Answer
1 | Black Body | Do I need an adapter plate to install the blackbody on the tripod? | No, there is a standard tripod screw hole at the bottom
2 | Black Body | How to connect blackbody and camera? | No need to connect, the blackbody can be used after power on, as a standard temperature source.
3 | Black Body | After turning on blackbody compensation, the measured temperature of the human body has risen significantly? | If the heat engine is guaranteed to be completed, it may be that the black body temperature control element is affected during transportation, resulting in a true temperature change (the accuracy will not change). The VCA resource mode of the thermal imaging device can be switched to temperature measurement, and the true temperature of the black body in the field of view of the camera can be measured. If the true temperature of the blackbody is 39.3 ° C, you can switch back to the human body temperature measurement mode and fill in the blackbody temperature at the blackbody compensation position as 39.3 ° C.
4 | Black Body | Can the temperature be measured outdoors after using a blackbody? | No, the outdoor environment affects the temperature of the measured skin surface, and has nothing to do with the black body calibration.
5 | Black Body | 1. An individual error occurs during measurement. The temperature measurement of other people is normal, but the temperature measurement of individual people is abnormal; 2. This situation is weakened when there is a black body, and it is more obvious when there is no black body; | 1. We measure the temperature of a person's forehead and the infrared radiation emitted by the person's forehead; 2. Therefore, not everyone's forehead temperature and body temperature are the same; at the same time, not everyone's emissivity is 0.98; 3. The black body can reduce the error, so when the black body exists, it can reduce the impact of this individual difference; 4. If such individual differences often occur, and customers are more concerned about this problem, it is recommended to deploy a black body in the system to reduce the occurrence of this problem

## Handheld
No. | Product Line | Question | Answer
1 | Handheld | Thermal imaging problem: 21B cannot connect to WIFI (both office and mobile phones are not allowed). Enter password-connection failed. There is currently only one device. Restart can not be resolved. | try use 21B as AP mode, let the computer connect.
2 | Handheld | Is only TP21B handheld temperature measurement in human body temperature measurement products supporting Hik-Thermal? | yes
3 | Handheld | Can TP21B be added to hik-connect？ | Hikconnect cansupport live-view after being added, but there is no event alarm for the time being (in the development plan)
4 | Handheld | Can TP21B work with black body together? | "Yes, but the function is limited. There are currently two ways to use: ① In Auto mode, that is, after fixing the relative position, place the black body on the screen for real-time calibration (need to draw the rule box), but the logic of the linkage alarm with the 4200 is more confusing when there is a rule box (whether the alarm is only Depends on the highest temperature of the screen, and has nothing to do with the rule box). This solution is not recommended at this time, and may be optimized in the future. ②Manual mode is equal to do a one-time calibration by using the blackbody. Within thirty minutes after the calibration, the measurement accuracy is increased from the original ± 0.5 ℃ to ± 0.3 ℃. "
5 | Handheld | Can TP31B support sound alarm? | No
6 | Handheld | What are the temperature measurement units of TP21B and TP31B? | Celsius (° C), Fahrenheit (° F), Kelvin (K)
7 | Handheld | Can the TP31B be connected externally or projected to the monitor? | No, it does not support access to all functional networks
8 | Handheld | How to display TP21B interface on the monitor? | TP21B is added to the computer 42 through WIFI, and the computer video output interface is connected to the monitor
9 | Handheld | Can the TP21B be directly wired to an external monitor? | Cannot
10 | Handheld | Can the handheld be connected to NVR to store video? | Yes, handheld camera access to the wifi of the network segment where nvr is located
11 | Handheld | Do I need an adapter plate to mount the handheld device on a tripod? | No, there is a standard tripod screw hole at the bottom
12 | Handheld | Can the TP21B be unplugged the battery and connected to the charging cable? | Yes, but not recommended. Bad operation will cause frequent power outages
13 | Handheld | What does Support audio intercom in TP21B spec mean? | Supports two-way intercom to connect to the event center (including Hik Thermal)
14 | Handheld | What does Offline Picture and Thermal Graphic mean in the camera setting of handheld products? | 1. Thermal Graphic is to save a snapshot; 2. Offline Picture can support secondary analysis, such as redrawing the area for temperature measurement, etc.
15 | Handheld | Hand-held TP21B, enter the WIFI password, after entering the password, click OK and nothing happens | 1. It is an occasional problem, and there is very little feedback at present, and it is difficult to confirm the reason for R & D; 2. It is recommended to initialize the device and then reconnect; 3. In addition, when the password is entered, choose Close, not Enter. Therefore, consider the password error caused by entering the character "Enter" in the password;
16 | Handheld | 1. Does the handheld support three-party batteries; 2. One charger can charge several batteries at the same time; 3. Can the device be used normally when recharging; | 1. The handheld product supports three parties' batteries; 2. One charger can charge 2 batteries at the same time; 3. The device can be used normally during the charging process;
17 | Handheld | Can the handheld product TP21B be connected to HCP? | After testing, I found: 1. HCP can access TP21B high-end handheld devices; 2. It can be previewed normally, and has been dual-channel with ordinary thermal camera; 3. The recording plan cannot be configured, but you can manually capture pictures and record videos on the Liveview interface; 4. Temperature and alarm can be configured, and the alarm information contains temperature information;
18 | Handheld | Can handheld devices connect to NVR? | 1. The handheld device supports access to Class I NVR; 2. However, it does not support the configuration of recording plans including various events such as event recording, scheduled recording, and 24-hour recording; 3. Only manual picture capture and recording are supported on the live-view interface;
19 | Handheld | Handheld device WIFI hotspot can access several mobile phones | In theory, up to 5 devices can be connected, but it is recommended not to exceed 3 devices.
20 | Handheld | Handheld device connected to Hik-Connect cannot receive high temperature alarm information | 1. The handheld device currently supports access to HIK-Connect; 2. However, Ezviz has not yet completed the high-temperature alarm function and is currently advancing. 3. If you are eager to use this function, you can Baidu [temperature measurement wizard], this app is developed by Ezviz;
21 | Handheld | Where is the Device Password of the handheld device set? I forgot my password | 1. The device password of the handheld device is set when the device is activated for the first time; 2. If you forget the password, you can only reset the device and set it after re-activation;
22 | Handheld | The picture of thermal channels are all orange | 1. It is necessary to cancel Above in Image Enhancement, that is, to cancel the function of high temperature highlighting; 2. The initial suspicion is that the ambient temperature in Indonesia is relatively high, and the ambient temperature exceeds the threshold temperature, resulting in the high temperature of the entire picture highlighting;
23 | Handheld | What is the capacity of the SD card that comes with the handheld device when it is shipped? | Economical 8G, high-end 16G
24 | Handheld | 1. In order to save storage space, does the handheld support capture screenshots? 2. Does SD card support storage coverage? 3. Can I use USB to export videos and screenshots? 4. Why is it written in the brochure that can be applied to the transformer substation scene? | 1. Those who do not support event capture; 2. The SD card does not support storage coverage; 3. Support USB export video and snapshot; 4. The customer downloads the TP31 brochure, TP31 and TP31B are different, TP31B measures the temperature of the human body, TP31 is not;
25 | Handheld | TP31 handheld device startup screen is pink | The camera movement is damaged and needs to be replaced free of charge
26 | Handheld | If the handheld temperature measurement tp21b is connected to nvr, if only the stream is connected to the NVR, is there temperature information superimposed on this stream? Can I see temperature information on the NVR? | Yes, TP21B itself has temperature information superimposed on the video stream.
27 | Handheld | Does TP21B high-end handheld support Onvif and ISAPI protocols? | Does not support Onvif, only supports ISAPI and SDK
28 | Handheld | Does the TP21B tripod connected to the standard screw interface need an adapter plate? | No need
29 | Handheld | Is the TP21B handheld add to 4200 by 8000 port? | The default is port 8000
30 | Handheld | What is the capacity of the expansion card supported by the handheld device? What are compatible brands? | Maximum support 128G; minimum configuration: class 4, read / write rate 5M / S; brand: Hikvision, Toshiba, SanDisk / Kingston are preferred, other brands do not guarantee compatibility.
31 | Handheld | Will the temperature information be automatically displayed after the TP21B is connected to the NVR? | Can

## NVR
No. | Product Line | Question | Answer
1 | NVR | When the thermal imaging camera is linked to the NVR, after the alarm pops up, the warning sound is delayed or does not sound | After repeated comparison tests, it was found that after the February version and the early March version were upgraded to the 0314 version, the front-end and back-end initialization problems can be solved At present, most cameras‘ problem can be solved, and a small part suspects that the NVR firmware vision is too old and incompatible. The NVR side will offer new firmware in late April. Firstly, upgarde the camera to 0314, then restart& restore factory. if problem not solved, wait for the new version of the baseline.
2 | NVR | Does I series NVR and Super Brain need to add the device to the 4200 if they want to receive the human body temperature alarm message? | No, just do it locally. However, it should be noted that now that the human body temperature measurement camera is added to the I series NVR, it may happen that only one channel is added, which can be added on PCweb
3 | NVR | Are non-I series NVRs compatible with thermal imaging equipment? For storage only. | Yes, you can see the temperature frame information in the storage video screen when you set the stream overlay.
4 | NVR | Does the i-series nvr currently support pass-through alarms? | Support transparent transmission, the current alarm and the previous industrial thermal imaging belong to the same temperature alarm category. The human body temperature measurement alarm, including the face snapshot and temperature data, requires the i-series NVR baseline version released at the end of April
5 | NVR | Add the thermal camera directly to the i series NVR at the NVR side, only one channel is added successfully | Log in to the IP of the i-series NVR from the computer and add the device from the WEB page to solve this issue.
6 | NVR | 1. The latest firmware version 0414 of the I series NVR will be released soon; 2. Human temperature display can be displayed in this version; 3. Sound alarm of temperature status; 4. Search and playback of temperature measurement results; 5. Temperature information transparent transmission alarm; 6. The event recording is still unsuccessful. Even if the event recording is selected, the event is still being recorded, but the NVR actually still records 24 hours; | R & D has been informed and they will improve afterwards
7 | NVR | Whether I series NVR alarm sound can be customized | The baseline does not support modification, it can only trigger the buzzer to sound
8 | NVR | When will the new NVR baseline firmware for body temperature measurement cameras be released | The beta version is already on ftp at / 01 All Materials of Temperature Screening Solution / 02 NVR / 05 Firmware / I Series NVR / The official baseline version will be released before mid-May
9 | NVR | Can the web interface of the I series NVR directly configure the camera's body temperature measurement configuration | Add via IP: not supported; via POE interface of NVR: can be configured with virtual host "
10 | NVR | Does the latest firmware of the I series NVR support event filtering based on the measured temperature | can support

## DeepinMind
No. | Product Line | Question | Answer
1 | DeepinMind | There is no temperature frame and temperature value information on the face image captured by DeepinMind | Face Capture >> Overlay & Capture>>enable Display Target Info. on Alarm Picture
2 | DeepinMind | The same person of the DeepinMind conducted multiple face captures and temperature measurements | Face Capture >> Advanced Configuration See if Capture Times is Limited as 1.
3 | DeepinMind | How to export the report in DeepinMind | 1. Only the information of Check in and Attandacne in DeepinMind can be exported in Excel format; 2. Other alarm information or personnel search does not support Excel format export, only photos and corresponding code information can be exported, but the exported photos contain temperature information;
4 | DeepinMind | After the upgrade of DeepinMind, there is no captured face picture | Upload captured face Image is not enabled on the camera side, and the fault is resolved after the function is enabled;
5 | DeepinMind | After upgrading the firmware of 77 DeepinMind, neither arming schdule nor audio warning can be enabled | Device upgrade failed, clear browser cookies, re-upgrade, restore device after upgrade
6 | DeepinMind | There are audio warnings in No mask linkage and abnormal linkage in DeepinMind. What is the difference between this and voice linkgae? | If you want to use the audio warning in no mask linkage and abnormal linkage, you must first enable the relevant options in voice linkage (Mask, Temp Normal, Temp Abnormal)
7 | DeepinMind | DeepinMind audible alarm failed | 1. DeepinMind supports audible alarms, but DeepinMind does not have a built-in speaker and requires an external speaker; 2. The camera also supports voice alarm, but only -P's camera comes with a built-in speaker, so no external speaker is required;
8 | DeepinMind | The DeepinMind connected to the ordinary camera of the temperature measurement scheme can display the face comparison on the UI interface. Is it detected whether it is with a mask? | Can
9 | DeepinMind | Is the front-end or back-end do the screenshot in the DeepinMind solution? | The working logic is as follows: the camera visible light channel captures the face + superimposed temperature information-the snapshot is given to the DeepinMind, the DeepinMind performs temperature classification, mask detection and face comparison

## 4200
No. | Product Line | Question | Answer
1 | 4200 | 3.2.0.10 version ivms-4200 export thermal imaging alarm screenshot only thermal imaging channel alarm screen, no optical channel screen | Feedback Zhan Hui, get the latest 4200 version 3.2.0.50 (ftp). The official website is still the old version.
2 | 4200 | The latest version of the baseline 4200 is connected to the thermal imaging camera Event search. When exporting the temperature measurement alarm event pictures, only the thermal imaging channel pictures are available. | Upgraded special temperature measurement version 4200 (supported after V3.2.0.50)
3 | 4200 | The temperature measurement Event picture storage and the selected local disk will generate the iVMS-4200alarmPicture folder, the generated file cannot be opened? | This file is a package of alarm pictures and cannot be opened directly. It needs to be exported in 4200
4 | 4200 | 4200 can only store thermal imaging channel pictures | use 4200 V3.2.0.102
5 | 4200 | The 4200 body temperature measurement version has no relevant interface | When installing the software, choose the plug-in AI DASHBOARD to install together
6 | 4200 | Human body temperature version 4200 AI DASHBOARD face recognition speed is very slow | By default, the face capture and temperature information will be uploaded when the face leaves the detection range
7 | 4200 | 4200 AI dashboard can't find the place to export data | The 4200 AI dashboard interface does not support exporting data, but you can search by event and export in Event;
8 | 4200 | What are the meanings of P1 and P2 on the TP21B interface | TP21B supports setting the maximum temperature point (Max.), The minimum temperature point (Min.) And the central temperature point (Cen.), And also supports user-defined point temperature measurement, P1 means the first custom point (Position) 1) And up to 10 custom temperature measurement points can be set;
9 | 4200 | Can the 4200 temperature measurement version offer a Minority language voice package? | Multilingual packages are expected to be given in 4.24
10 | 4200 | Can the 4200's AI dashboard display face comparison and temperature measurement at the same time? | can
11 | 4200 | What is the latest version of 4200 in FTP? | 3.2.50.2, the latest customized version released for Haikang's full line of human body temperature measurement products

## Other
No. | Product Line | Question | Answer
1 | other | The usage scene of the device is used in entrance, but we recommend that people enter the scene from the outside, and check it after 5 minutes. In this case, is it impossible to recommend this solution at the entrances of some shopping malls, airports and other scenes? | If the temperature difference between the indoor and outdoor is not too large and the impact is small, it can be recommended to use a drainage channel and other solutions to try to make the room as much as possible before entering the temperature measurement environment (preferably 3-5min); If the temperature difference is more obvious and the temperature is relatively low / high (no wind in the room), manual compensation can be considered for the barrel machine and hemisphere, but there is no relevant setting for handheld product.
2 | other | Can hik-thermal app receive temperature alarm? | No, no event center in hik-thermal (in development plan)
3 | other
4 | other | Does Milestone support the reception of temperature alarms from body temperature measuring cameras | Yes. The latest plugin is already supported.
5 | other | Why the device cannot be added to Hik-Connect | Network >> Advanced Settings >> Platform Access>>configure Hik-connect
6 | other | Why the human body temperature measuring device added to Hik-Connect cannot receive the alarm | It is recommended to add it by scanning the QR code or filling in the serial number; adding directly via IP will not receive the alarm
7 | other | 1. What do these icons mean? 2. Why the temperature threshold is cleared at the click and the camera keeps alarming | 1. Respectively measure the center temperature, the highest temperature point, the lowest temperature point, and the frame temperature; 2. This is a bug of the current product. The temperature threshold will be cleared when clicked, and the camera will always alarm; 3. Currently solving this defect;
8 | other | Hikconnect and human body temperature measurement equipment alarm logic | Currently: the same channel, only one event will be uploaded every minute

## HCP
No. | Product Line | Question | Answer
1 | HCP | How about compatibility with HikCentral? Can the temperature measurement analysis report be generated? | HCP has just released a customized version for human body temperature measuring, which supports general functions such as preview alarm, but has not yet been released.
2 | HCP | HCP + 2TD1217B, when there is no DeepinMind, will there be real-time temperature records here? | 1. The Body Temperature module is intended for temporary scenarios. In this scenario, there is no DeepinMind to participate in; 2. This interface will display the real-time temperature; 3. However, this interface is not a Live View interface, it will only display a picture of the person who triggered the threshold; 4. There are two thresholds on the HCP. One threshold is used to filter out extreme cases of abnormally low body temperature (such as fake faces), and the other threshold is used to filter out abnormal body temperatures (such as fever);
3 | HCP | After setting the temperature on the HCP temperature measurement module Body Temperature, only the people with abnormal high temperature will be displayed, and the people with normal body temperature will no longer be displayed. | 1. There are two thresholds on the HCP. One threshold (Thermometry Threshold) is used to filter out extreme cases of abnormally low body temperature (such as fake faces), and the other threshold (Alarm Threshold) is used to filter out abnormalities (such as fever); 2. Thermometry Threhold will be synchronized to the alarm value and pre-alarm value of the device after the setting is completed, and the values of these two parameters remain the same; 3. Set the Thermometry threshold to about 35, which is used to choose less extreme low-temperature conditions; set the Alarm Threshold to the high-temperature alarm threshold, which will display the normal body temperature and the body temperature of people with abnormal temperature;
4 | HCP | HCP temperature measurement module Body Temperature, there is no way to link the camera | 1. The HCP temperature measurement module (Body Temperature) is for the temporary deployment of the scene, that is, the thermal imaging camera stand alone; 2. This module does not support indirect association via DeepinMind, and can only be directly connected to HCP;